Recommendations

Generic statement before any recommendations:

*Statement on Dose Optimization Recommendation*

Please note that your CTDIvol and/or DLP is higher than the recommended DRL for dose optimization. If you want to reduce radiation doses, you may consider the following:

Pediatric head:

1. As a first step towards radiation dose optimization, please consider reducing the radiation dose just below the DRL for the specified CT exam type and patient age.
2. Please ensure proper patient centering.
3. Avoid using 140 kV
4. Consider reducing tube current. Please remember that tube current has a linear relation with radiation dose. A 25% reduction in tube current will lead to 25% decrease in radiation dose if all other scan factors are constant.
5. When available, please consider using iterative reconstruction (IR) or deep learning (DL) reconstruction techniques to reduce image noise. Choice of IR strength varies based on the scanner and vendor (typically a strength of 2-3 (Siemens/Philips), mild-moderate (Canon) or 30-40% (GE) is sufficient for most CT protocols).

Pediatric Chest

1. Please consider reducing the radiation dose just below the DRL for the specified CT exam type and patient size.
2. Please ensure proper patient centering.
3. For chest CT, please make sure to keep a fast acquisition time (with higher pitch and/or fast rotation time) to minimize motion artifacts. It is important to avoid or minimize motion artifacts to avoid loss of diagnostic information and unnecessary repeat scanning.
4. For patients <50 kg, 80 KV is typically sufficient. For patients 50-80 kg, 100 kV is generally sufficient for chest CT.
5. You can also consider decreasing tube current for reducing radiation dose. Please remember that tube current has a linear relation with radiation dose. A 25% reduction in tube current will lead to 25% decrease in radiation dose if all other scan factors are constant.
6. Most chest CT must be performed with a single-phase chest CT protocol. Perform either non-contrast chest CT or post-contrast CT based on clinical need.
7. When available, please consider using iterative reconstruction (IR) or deep learning (DL) reconstruction techniques to reduce image noise. Choice of IR strength varies based on the scanner and vendor (typically a strength of 2-3 (Siemens/Philips), mild-moderate (Canon) or 30-40% (GE) is sufficient for most CT protocols).

Pediatric abdomen:

1. Please consider reducing the radiation dose just below the DRL for the specified CT exam type and patient size.
2. Please ensure proper patient centering.
3. Avoid use of 140 kV for abdomen CT. You can use 100-120 kV for smaller children (<50 kg). Or you can use automatic tube potential selection technique.
4. You can consider decreasing tube current for reducing radiation dose. Please remember that tube current has a linear relation with radiation dose. A 25% reduction in tube current will lead to 25% decrease in radiation dose if all other scan factors are constant.
5. Most routine pediatric abdomen CT exams should be done with a single-phase CT.
6. When performing multi-phase abdomen CT, ensure proper clinical justification (such as for liver or renal lesion evaluation or hematuria).
7. For multiphase CT, often the scan length for some phases can be limited to the organ of main clinical interest. For example, limiting arterial phase CT to liver for liver lesions. Likewise, eliminating certain phases such as non-contrast phase for liver CT can also help reduce dose.
8. When available, please consider using iterative reconstruction (IR) or deep learning (DL) reconstruction techniques to reduce image noise. Choice of IR strength varies based on the scanner and vendor (typically a strength of 2-3 (Siemens/Philips), mild-moderate (Canon) or 30-40% (GE) is sufficient for most CT protocols).

Adult head:

1. Please consider reducing the radiation dose just below the DRL for the specified CT exam type.
2. Please ensure proper patient centering.
3. Avoid using 140 kV
4. Consider reducing tube current. Please remember that tube current has a linear relation with radiation dose. A 25% reduction in tube current will lead to 25% decrease in radiation dose if all other scan factors are constant.
5. When available, please consider using iterative reconstruction (IR) or deep learning (DL) reconstruction techniques to reduce image noise. Choice of IR strength varies based on the scanner and vendor (typically a strength of 2-3 (Siemens/Philips), mild-moderate (Canon) or 30-40% (GE) is sufficient for most CT protocols).

Adult Chest:

1. **Radiation doses greater than the DRLs are acceptable for large patients.**
2. For non-obese patients (typically < 70-80 kg), please consider reducing the radiation dose just below the DRL for the specified CT exam type.
3. Please ensure proper patient centering.
4. Please keep a fast acquisition time (with higher pitch and/or fast rotation time) to minimize motion artifacts. It is important to avoid or minimize motion artifacts to avoid loss of diagnostic information and unnecessary repeat scanning.
5. For patients <50 kg, 80 KV is typically sufficient. For patients 50-80 kg, 100 kV is generally sufficient for chest CT.
6. For some patients, you can reduce tube current for reducing radiation dose. Please remember that tube current has a linear relation with radiation dose. A 25% reduction in tube current will lead to 25% decrease in radiation dose if all other scan factors are constant.
7. Most chest CT must be performed with a single-phase chest CT protocol. Perform either non-contrast chest CT or post-contrast CT based on clinical need.
8. For dual-phase chest CT (such as in inspiration and expiration) in patients with interstitial lung diseases or tracheal collapsibility, please perform expiration phase CT with low radiation dose (at least 50% below the inspiration phase dose).
9. When available, please consider using iterative reconstruction (IR) or deep learning (DL) reconstruction techniques to reduce image noise. Choice of IR strength varies based on the scanner and vendor (typically a strength of 2-3 (Siemens/Philips), mild-moderate (Canon) or 30-40% (GE) is sufficient for most CT protocols).

Adult abdomen:

1. **Radiation doses greater than the DRLs are acceptable for large patients.**
2. For non-obese patients (typically < 70-80 kg), please consider reducing the radiation dose just below the DRL for the specified CT exam type.
3. Please ensure proper patient centering.
4. Consider decreasing tube current for reducing radiation dose. Please remember that tube current has a linear relation with radiation dose. A 25% reduction in tube current will lead to 25% decrease in radiation dose if all other scan factors are constant.
5. Avoid use of 140 kV for abdomen CT for non-obese patients.
6. Most routine abdomen CT exams should be done with a single-phase CT.
7. When performing multi-phase abdomen CT, ensure proper clinical justification (such as for liver or renal lesion evaluation or hematuria).
8. For multiphase CT, often the scan length for some phases can be limited to the organ of main clinical interest. For example, limiting arterial phase CT to liver for liver lesions. Likewise, eliminating certain phases such as non-contrast phase for liver CT can also help reduce dose.
9. When available, please consider using iterative reconstruction (IR) or deep learning (DL) reconstruction techniques to reduce image noise. Choice of IR strength varies based on the scanner and vendor (typically a strength of 2-3 (Siemens/Philips), mild-moderate (Canon) or 30-40% (GE) is sufficient for most CT protocols).